Measurement Report

| Subject of the exercise: | <Title> (Exercise <No>.) |
| --- | --- |
| Date: | <year>. <month>. <day> |
| Students name: | <name 1> <name 2> <name 3> |
| Course and group No. | Course: P2, <Group No.> |
| Supervisors: | <name 1>, <name 2> |

Equipment in use, devices under test

| Precision Magnetics Analyzer | Wayne Kerr 3260B | <Unit no.> |
| --- | --- | --- |
| Test panel | VIK– 01–07 | <Unit no.> |

Measurement Tasks

- Measurement of ferromagnetic properties of materials

Technical data of core:

- Type: TDK H5A

- Material: ferrite

- Shape: toroidal

- Dimensions: D = 68 mm, d = 44 mm, h = 13,5 mm

- Bandwidth: max. 0,2 MHz

- Initial permeability, μi : 3300   -0….+40%

- Max. flux density, Bm: 410 mT

- Inductivity factor, AL: 4300±25% nH

  - Measure the coil impedance at 1V in the range of 100 Hz…500 kHz!

<your comments>

  - Measure the resonance frequency and the self-capacitance of the coil!

<your comments>

  - What is the max. measurement frequency if the systematic error caused by the resonance is not higher than 1%?

<your comments>

  - Measure the coil impedance at 150 Hz in the range of 1 mV…10 V using the Graph-mode of the Analyzer!

<your comments>

  - Calculate μr(Bm) and AL(Bm) from the Z(U) characteristics!

<your comments>

- Measurement of model parameters of transformers

Technical data of DUT No Tr-1 and Tr-2:

- Type: TDK H5A

- Core material: ferrite

- Shape: toroidal

- Dimensions: D = 68 mm, d = 44 mm, h = 13,5 mm

- Frequency: max. 0,2 MHz

- Initial permeability, μi : 3300   -0….+40%

- Max. flux density, Bm: 410 mT

- Inductivity factor, AL: 4300±25% nH

- Number or primary turns, Np: 110

- Number or secondary turns, Ns: 110

- Cross section of the copper wire, ACu: 0,05 mm2

- Diameter of isolated wire dv: 0,6 mm

- Coupling: tight coupling (Tr-1), loose coupling (Tr-2)

  - Measure the model parameters at Ueff = 5V and f = 1 kHz! Measure the dc parameters! What are the differences between the two transformers?

<your comments>

  - Estimate the measurement uncertainties of the copper resistances and that of the magnetizing inductivity!

<your comments>

  - Measure the primary impedance Z(U, f) of the unloaded transformer!

<your comments>

- In-circuit measurement on RC-network

Technical data of the RC-components:

- R1 = 100 Ω ±0,1%	
			R2 = R4 = 1 kΩ ±1%	
			R5 = 10 kΩ ±2%	
			C1 = 1 μF ±5%	
			C2 = 100 nF ±5%

Figure 7–1. Schematic of low-pass filter to be measured

  - Measure all of the RC components using in-circuit technique! Are the measured values within the specified tolerance bands?

<your comments>

- Measurement on Telecom transformer

Technical date of Tr-3:

- Type: PCM-40

- Core material: M30 Permalloy

- Turn ratio, Ns/Np: 408:384

- Frequency band: 200 Hz…220 kHz

- Insertion loss, IL: max. 0,1 dB

- Max. voltage level: +10 dB

- Terminating impedance: 600 Ω

  - Measure the turn ratio of the Telecom transformer!

<your comments>

  - Measure the Insertion Loss and Reflection Loss of the Telecom transformer at given frequency using direct coupling!

<your comments>

  - Measure the Insertion Loss and Reflection Loss of the Telecom transformer at given frequency using direct coupling and RC-damping!

<your comments>

  - Measure the Insertion Loss and Reflection Loss of the Telecom transformer at given frequency using capacitance coupling!

<your comments>